## Slide 1
Detector Efficiency and Peak Fitting
1
What you need to know:
  How semiconductor detectors work
  How to determine the detection efficiency
  How to find nuclear information
  How to fit peaks with GRSISort
  How to run a script in ROOT/GRSISort
What you need to do: The detection efficiency calibration of real germanium detector arrays!
  Install GRSISort
  Fit gamma-ray photopeaks with GRSISort
  Edit and run the RootEffi script
  Submit screenshots of your fitted peaks, edited RootEffi script and efficiency curve 

## Slide 2
Germanium Detector Arrays
2
Soccer Ball,
iThemba LABS,
Cape Town,
South Africa
52 crystals
TIGRESS,
TRIUMF, Vancouver,
Canada
56 crystals
GRIFFIN,
TRIUMF,
Vancouver,
Canada
64 crystals
MINIBALL,
CERN,
Geneva,
Switzerland
24 crystals

## Slide 3
Configurations of Germanium Detectors
3
(13 clovers, 4 segmented and 9 not segmented)
TIGRESS,
TRIUMF, Vancouver,
Canada
56 crystals
GRIFFIN,
TRIUMF,
Vancouver,
Canada
64 crystals
MINIBALL,
CERN,
Geneva,
Switzerland
24 crystals
4 tapered Coaxial crystals
Cold finger
LN2 reservoir
Preamp, HV filter
Closed-end coaxial right cylinders 
(16 clovers, not segmented)
Soccer Ball,
iThemba LABS,
Cape Town,
South Africa
52 crystals
(8 triple clusters, 6-fold segmented)
(14 clovers, 8-fold segmented)

## Slide 4
How Semiconductor Detectors Work
4
An external field creates an area depleted of free charge carriers.
W.R. Leo, Techniques for Nuclear and Particle Physics Experiments (1994).
G.F. Knoll, Radiation Detection and Measurement (1989).
K.S. Krane, Introductory Nuclear Physics (1988).
Radiation interacts with the crystal and produces of electron-hole pairs.
The electrons and holes drift towards electrodes and the electric pulse is amplified and processed.

## Slide 5
Semiconductors vs. Scintillators
5
Material Property | NaI(Tl) | Germanium | Silicon
Type | Scintillator | Semiconductor | Semiconductor
Cooling | No | ~ 95 K | No
Density (g/cm2) | 3.7 | 5.3 | 2.3
Band gap energy (eV) | 20 | 2.9 | 3.8
Energy resolution at 1332 keV (keV) | 13 | 1.9 | 

## Slide 6
How to Determine the Detection Efficiency
6
Area of the photopeak
Data collection time
Activity of the source
Absolute intensity of the gamma ray

## Slide 7
How to Find Nuclear Information
7
Lund/LBNL Nuclear Data Search: http://nucleardata.nuclear.lu.se/toi/
National Nuclear Data Center: https://www.nndc.bnl.gov/